## List
List of BGI Standard Commercial Cyclones |  |  |  | Rev.12, Feb, 2014 | 2014
Type No. | Name | D50-μm | Q - lpm | Sharpness | Calculator
VSCC 2.946 | EPA Equivalent | 2.5 | 16.67 | 1.157 | Y
SCC 1.062 | Triplex | 1/2.5/4 | 3.5/1.5/1 | 1.21 | Y
SCC 1.829 | Photometer/Environmental Sampler | 2.5 | 5 | 1.23 est | Y
SCC 2.654 | Photometer/Environmental Sampler | 2.5 | 10 | 1.22 est | Y
SCC 2.229 | PM1 | 1 | 16.67 | 1.17 | Y
SCC 3.495 | PM2.5 | 2.5 | 16.67 | 1.19 | Y
GK 2.05 | KTL | 2.5 | 4 | 1.32 | Y
GK 2.05 SH | KTL | 2.5 | 3.5 | 1.32 | Y
GK 2.69 | Resp./Thoraccic | 4/10 | 4.2/1.6 |  | Y
BGI 4 | H-D | 4 | 2.2 | 1.22/1.22 est | Y
SCC 0.732 | Photometer/Environmental Sampler | 1 | 2 | 1.19 est | Y
SCC 1.197 | Photometer/Environmental Sampler | 1/2.5 | 5/2.27 | 1.17/1.25 est | Y
SCC 2.354 | Photometer/Environmental Sampler | 1/2.5 | 17.3/8 |  | Y
SCC 0.695 | Diesel soot | 0.8 | 2.2 | 1.2 | Y
GK 4.162 | Rascal Respirable | 4 | 9.5 | 1.27 | Y
The purpose of this list is to present a calculator for each of the BGI cyclones for which there is a mathematical model. It should be
understood that these "calculators" are derived from the model which is deduced from actual test results. Where a cyclone has
been tested at a particular set of conditions, the information is reliable. Extrapolations are assumed to be reliable.
The main purpose of the calculator is to assist investigators who have operated at a flow rate for which there is no data, to
appreciate their results. The secondary purpose of the calculator is to select a set of operating parameters for a specific study.
All pressure drop data was collected as a separate study.
All test data was taken at Engineering standard conditions.
The publication and public availability of the information contained in these documents does not place BGI under obligation as to
the correctness of the information nor any consequences resulting from the usage thereof .

## SCC 2.354
SCC 2.354
2 | 12.495
3 | 7.804
4 | 5.588
5 | 4.312
6 | 3.49
7 | 2.918
8 | 2.499
9 | 2.179
10 | 1.929
11 | 1.727
12 | 1.561
13 | 1.422
14 | 1.305
15 | 1.204
16 | 1.117
17 | 1.042
18 | 0.975
Work Area
LPM | Q
8 | 2.4988
Enter in Red
Answer in Blue
Work Area
SCC 2.354 |  |  |  | Enter in RED
Mfg: BGI |  |  |  | Answer in BLUE
Dist: Sunset Labs |  |  |  | Q-lpm | DP
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | 10 | 2.9328
Room BP: | 744.6 | mm of Hg
RoomTemp: | 19.3 | C
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1D | 121.3 | 19.2 | 1434.5854 | 0.61 | 0.59821
| 170.4 | 19.3 | 2034.5611 | 1.28 | 1.2553
| 222.6 | 19.3 | 2671.8954 | 2.345 | 2.2997
| 272.2 | 19.3 | 3277.4851 | 3.7 | 3.6285
| 320.9 | 19.3 | 3872.0862 | 5.2 | 5.0995
| 358.3 | 19.3 | 4328.7204 | 6.85 | 6.7176
| 413.2 | 19.3 | 4999.0203 | 9.6 | 9.4144
| 459.3 | 19.3 | 5561.8769 | 12.2 | 11.9641
| 498.7 | 19.3 | 6042.93 | 14.8 | 14.5138
Room BP: | 743.4 | mm of Hg
RoomTemp: | 19.3 | C
#2D | 201.9 | 19.6 | #DIV/0! | 0.12 | 0.11768
| 280 | 19.6 | #DIV/0! | 0.215 | 0.21084
| 373.7 | 19.6 | #DIV/0! | 0.385 | 0.37756

## VSCC 2.946
Lpm | D50 |  |  |  |  | HSE test data
2 | 19.768 |  |  |  |  | 18 | 2.36
4 | 10.057 |  |  |  |  | 15 | 2.758
6 | 6.773 |  |  |  |  | 11.4 | 3.57
8 | 5.116 |  |  |  |  | 15.7 | 2.66
10 | 4.116 |  |  |  |  | 18.7 | 2.295
12 | 3.446
14 | 2.965
16 | 2.603
18 | 2.32
20 | 2.094
D50= | 2.5009
Lpm= | 16.67
Work Area
Enter in Red
Answer in Blue
Q-lpm | D50
47 | 0.99847
2008-07-31 | Operator: | rg |  | Work Area
VSCC 2.946 | S/N 110308-5 |  |  | Enter in RED
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Answer in BLUE
Room BP: | 751.3 | mm of Hg |  | Q-lpm | SP
RoomTemp: | 22.5 | C |  | 47 | 20.6637
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1 | 128 | 22.2 | 4.9439 | 0.195 | 0.19123
| 174.95 | 22.3 | 6.8327 | 0.35 | 0.34323
| 229 | 22.3 | 9.0052 | 0.625 | 0.61292
| 279.45 | 22.2 | 11.0292 | 0.925 | 0.90712
| 319.8 | 22.3 | 12.6547 | 1.225 | 1.2013
| 376.2 | 22.2 | 14.9166 | 1.75 | 1.7162
| 419.5 | 22.4 | 16.6677 | 2.195 | 2.1526
| 457.2 | 22.4 | 18.1835 | 2.685 | 2.6331
| 501.3 | 22 | 19.9296 | 3.27 | 3.2068
| 549.2 | 22.1 | 21.8603 | 4.025 | 3.9472
| 621.15 | 22.2 | 24.7586 | 5.2 | 5.0995
| 708.8 | 22.2 | 28.2804 | 6.95 | 6.8156
#2 | 203.6 | 21.8 | 7.9707 | 0.035 | 0.034323
| 296.5 | 21.8 | 11.6984 | 0.09 | 0.08826
| 359.4 | 21.8 | 14.2223 | 0.115 | 0.11278

## SCC 0.695
SCC 0.695 no calibration chamber
Q | D 50
lpm | µm
1.5 | 1.25
1.6 | 1.16
1.7 | 1.081
1.8 | 1.011
1.9 | 0.95
2 | 0.895
2.1 | 0.846
2.2 | 0.801
2.3 | 0.761
2.4 | 0.724
2.5 | 0.691
Work Area
Enter in Red
Return in Blue
Q-lpm | D50
0.82 | 2.5203
2010-04-05 | Operator: | rg
SCC 0.695 |  |  |  | Work Area
Qs corrected to: BP=760 mm of Hg, T= 20C |  |  |  | Enter in RED
Room BP: | 758.17 | mm of Hg |  | Answer in BLUE
RoomTemp: | 22.2 | C |  | Q-lpm | SP
|  |  |  | 2.2 | 10.6517
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
Room BP: | 762.1 | mm of Hg
RoomTemp: | 22.5 | C
#2D | 152.4 | 22.2 | 1573.3452 | 4.8 | 4.7072
| 170.1 | 22.2 | 1761.2077 | 6.3 | 6.1782
| 189.91 | 22.2 | 1971.4651 | 8.2 | 8.0415
| 209.43 | 22.3 | 2179.3821 | 10.3 | 10.1008
| 230.22 | 22.3 | 2400.1157 | 13.1 | 12.8467

## SCC 1.062
SCC1.062 no calibration chamber
Q | D50
3320 Nov
4 | 0.86
1.8 | 1.97
4 | 0.85
1.2 | 3.1 | 3310 Nov
1.8 | 2.1
1.2 | 3.2
1.5 | 2.45
1.6 | 2.45
1 | 4.2
1.1 | 3.75
3 | 1.14 | 3320 Jan
4 | 0.86
3.5 | 1.01
Work Area
Enter in Red
Answer in Blue
Q-lpm | D50
8 | 0.39209
2008-08-27 | Operator: | rg |  | Work Area
SCC1.062 |  |  |  | Enter in RED
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Answer in BLUE
Room BP: | 762.1 | mm of Hg |  | Q-lpm | SP
RoomTemp: | 22.5 | C |  | 3.5 | 8.1881
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
#2D | 84.45 | 22.3 | 0.8512 | 0.29 | 0.28439
| 150.8 | 22.2 | 1.5541 | 1.01 | 0.99047
| 205.9 | 22.2 | 2.1381 | 2.04 | 2.0006
| 300.1 | 22.3 | 3.1375 | 4.8 | 4.7072
| 388.55 | 22.4 | 4.0766 | 8.7 | 8.5318

## SCC 1.829
Q | D50 | SCC 1.829 | rg | 2008-02-25
2 | 7.243
2.5 | 5.59
3 | 4.523
3.5 | 3.782
4 | 3.239
4.5 | 2.825
5 | 2.51
5.5 | 2.238
6 | 2.023
Work Area
Enter in RED
Answer in BLUE
Q-lpm | D50
8 | 1.4504
2009-01-20 |  |  |  | Work Area
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Enter in RED
Room BP: | 750 | mm of Hg |  | Answer in BLUE
RoomTemp: | 19.6 | C |  | Q-lpm | SP
|  |  |  | 10 | 7.0889
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1D | 179.9 | 20.2 | 7.1877 | 3.42 | 3.3539
| 223.1 | 20.2 | 8.9506 | 5.4 | 5.2956
| 274.3 | 20.2 | 11.0399 | 8.9 | 8.7279
Room BP: | 750 | mm of Hg
RoomTemp: | 19.6 | C
#2D | 148.7 | 20.1 | 1.5455 | 0.13 | 0.12749
| 254.6 | 20.1 | 2.6779 | 0.37 | 0.36285
| 349.9 | 20.1 | 3.6969 | 0.78 | 0.76492
| 461.5 | 20.1 | 4.8902 | 1.435 | 1.4073
| 599.4 | 20.1 | 6.3647 | 2.61 | 2.5595

## SCC 2.654
Lpm | D50
5 | 5.548
6 | 4.489
7 | 3.754
8 | 3.215
9 | 2.804
10 | 2.481
11 | 2.221
12 | 2.008
13 | 1.829
14 | 1.679
15 | 1.549
Work area
Enter in Red
Answer in Blue
Q-lpm | D50
10 | 2.481
2009-01-20 | Operator: | rg
SCC 2.654 |  |  |  | Work Area
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Enter in RED
Room BP: | 751.8 | mm of Hg |  | Answer in BLUE
RoomTemp: | 19.6 | C |  | Q-lpm | SP
|  |  |  | 10 | 1.5301
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1D | 176.8 | 20 | 7.0395 | 0.71 | 0.69627
| 225.7 | 20.1 | 9.0319 | 1.24 | 1.216
| 274.7 | 20.1 | 11.026 | 1.94 | 1.9025
| 319.7 | 20.1 | 12.8573 | 2.43 | 2.383
| 375 | 20 | 15.1026 | 3.9 | 3.8246
| 429.8 | 19.9 | 17.326 | 5.1 | 5.0014
| 528.7 | 19.9 | 21.3481 | 8.4 | 8.2376
Room BP: | 751.8 | mm of Hg
RoomTemp: | 19.6 | C
#2D | 353.9 | 20 | 3.7294 | 0.18 | 0.17652
| 458.9 | 20 | 4.8491 | 0.3 | 0.2942

## SCC 2.229
SCC D50 vs. Lpm
Lpm | D50
10 | 1.806
11 | 1.618
12 | 1.464
13 | 1.335
14 | 1.225
15 | 1.132
16 | 1.051
17 | 0.98
18 | 0.917
19 | 0.862
20 | 0.813
Work area
Enter in Red
Answer in Blue
Q | D50
16.67 | 1.0022
2008-08-04 | Operator: | rg
SCC 3.495 |  |  |  | Work Area
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Enter in RED
Room BP: | 756.1 | mm of Hg |  | Answer in BLUE
RoomTemp: | 22.5 | C |  | Q-lpm | SP
|  |  |  | 16.67 | 10.1
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1 | 124.8 | 22.6 | 4.7913 | 0.615 | 0.60311
| 175.9 | 22.6 | 6.8342 | 1.32 | 1.2945
| 223.9 | 22.6 | 8.7532 | 2.31 | 2.2653
| 274.9 | 22.6 | 10.7921 | 3.735 | 3.6628
| 320.5 | 22.4 | 12.6066 | 5.2 | 5.0995
| 374.8 | 22.6 | 14.786 | 7.5 | 7.355
| 418.1 | 22.6 | 16.5171 | 9.7 | 9.5125
| 459.35 | 22.2 | 18.1416 | 12.2 | 11.9641
| 499.75 | 22.1 | 19.7479 | 15 | 14.71
| 549.3 | 22.1 | 21.7255 | 18.5 | 18.1423
| 620.8 | 22 | 24.5708 | 24.5 | 24.0263
| 692.2 | 22.1 | 27.4288 | 30.2 | 29.6161
#2 | 207 | 22.6 | 8.0775 | 0.135 | 0.13239
| 296.9 | 22.6 | 11.6716 | 0.275 | 0.26968
| 357 | 22.6 | 14.0744 | 0.42 | 0.41188

## SCC 3.495
SCC D50 vs. Lpm
Lpm | D50
10 | 4.525
11 | 4.051
12 | 3.622
13 | 3.337
14 | 3.062
15 | 2.826
16 | 2.622
17 | 2.444
18 | 2.287
19 | 2.148
20 | 2.024
Work area
Enter in Red
Answer in Blue
Q | D50
36.7 | 1.0033
2008-08-04 | Operator: | rg |  | Work Area
SCC 3.495 |  |  |  | Enter in RED
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Answer in BLUE
Room BP: | 756.1 | mm of Hg |  | Q-lpm | SP
RoomTemp: | 22.5 | C |  | 11 | 0.91931
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1 | 125 | 21.9 | 4.7879 | 0.17 | 0.16671
| 176 | 21.9 | 6.822 | 0.34 | 0.33343
| 224.3 | 22.3 | 8.7603 | 0.56 | 0.54917
| 275.9 | 22.2 | 10.8174 | 0.87 | 0.85318
| 320 | 22.3 | 12.5824 | 1.195 | 1.1719
| 376.2 | 22.2 | 14.8219 | 1.69 | 1.6573
| 419.6 | 22.4 | 16.5658 | 2.15 | 2.1084
| 459.1 | 22.4 | 18.1439 | 2.6 | 2.5497
| 499.45 | 22 | 19.7292 | 3.25 | 3.1872
| 549.8 | 22.1 | 21.7455 | 3.99 | 3.9129
| 620.7 | 22.2 | 24.5835 | 5.1 | 5.0014
| 691.25 | 22.2 | 27.4002 | 6.3 | 6.1782
#2 | 205.4 | 21.8 | 7.9919 | 0.045 | 0.04413
| 300.2 | 21.8 | 11.7716 | 0.095 | 0.093163
| 356.3 | 21.8 | 14.0084 | 0.12 | 0.11768

## GK 2.05
KTL Cyclone. D50 vs. Q. Rg - 22 Mar 01
LPM | D50
1.5 | 7.668
2 | 5.519
2.5 | 4.277
3 | 3.472
3.5 | 2.911
4 | 2.499
4.5 | 2.184
5 | 1.937
Work area
Enter in Red
Answer in Blue
Q | D50
8 | 1.1316
2008-08-11 | Operator: | rg |  | Work Area
GK 2.05 |  |  |  | Enter in RED
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Answer in BLUE
Room BP: | 754.2 | mm of Hg |  | Q-lpm | SP
RoomTemp: | 22.5 | C |  | 10 | 8.6084
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1 | 123.9 | 21.3 | 4.7463 | 1.41 | 1.3827
| 175.3 | 21.3 | 6.7974 | 3.49 | 3.4225
| 222.8 | 21.2 | 8.6898 | 6 | 5.884
| 275.6 | 21.2 | 10.796 | 10.4 | 10.1989
| 319.9 | 21.2 | 12.5631 | 15 | 14.71
Room BP: | 754.7 | mm of Hg
RoomTemp: | 21 | C
#2 | 119.8 | 21 | 1.3135 | 0.064 | 0.062763
| 156.1 | 21 | 1.7258 | 0.14 | 0.13729
| 203.5 | 21.1 | 2.265 | 0.245 | 0.24026
| 285.05 | 21.2 | 3.1927 | 0.51 | 0.50014
| 356 | 21.2 | 3.9992 | 0.89 | 0.87279

## GK 2.05 SH
KTL Cyclone. D50 vs. Q. Rg - 22 Mar 01
LPM | D50
1.5 | 5.519
2 | 4.277
2.5 | 3.472
3 | 2.911
3.5 | 2.5
4 | 2.18
4.5 | 1.94
Work area
Enter in Red
Answer in Blue
Q | D50
8.5 | 1.0672
2008-08-11 | Operator: | rg |  | Work Area
GK 2.05 SH |  |  |  | Enter in RED
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Answer in BLUE
Room BP: | 755.2 | mm of Hg |  | Q-lpm | SP
RoomTemp: | 21 | C |  | 10 | 13.1649
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1 | 124.75 | 20.9 | 4.7674 | 2.48 | 2.432
| 176.6 | 21.2 | 6.8378 | 5.3 | 5.1975
| 219.5 | 21.2 | 8.5468 | 9.1 | 8.9241
| 274.8 | 21.2 | 10.7498 | 15.6 | 15.2984
| 328 | 21.1 | 12.8648 | 23.7 | 23.2418
Room BP: | 755.1 | mm of Hg
RoomTemp: | 21 | C
#2 | 112.1 | 20.9 | 1.2249 | 0.125 | 0.12258
| 154 | 20.8 | 1.6999 | 0.205 | 0.20104
| 204.7 | 20.8 | 2.2751 | 0.44 | 0.43149
| 282.6 | 20.8 | 3.1589 | 0.89 | 0.87279
| 353.4 | 20.8 | 3.9622 | 1.6 | 1.5691

## GK 2.69
GK 2.69 Cyclone. D50 vs. Q |  |  | rg | 2007-08-06
Lpm | D50
1.6 | 10
2.7 | 6.22
4.2 | 4
5 | 3.47
6 | 2.82
7 | 2.36
8 | 2.03
9 | 1.77
10 | 1.57
11 | 1.41
12 | 1.27
13 | 1.16
14 | 1.07
15 | 0.988
Work Area
Enter in Red
Answer in Blue
Q | D50
6 | 2.8149
2008-08-08 | Operator: | rg |  | Work Area
GK 2.69 |  |  |  | Enter in RED
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Answer in BLUE
Room BP: | 751.6 | mm of Hg |  | Q-lpm | SP
RoomTemp: | 21.5 | C |  | 9 | 2.6067
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1 | 126.8 | 20.9 | 4.8722 | 0.66 | 0.64724
| 174.8 | 20.9 |  | 0.905 | 0.8875
| 224.4 | 20.9 | 8.775 | 2.53 | 2.4811
| 274.6 | 20.9 | 10.7823 | 4.05 | 3.9717
| 319.9 | 20.9 | 12.5937 | 5.7 | 5.5898
| 375 | 20.9 | 14.797 | 8.6 | 8.4337
| 416.9 | 20.9 | 16.4725 | 11.2 | 10.9834
Room BP: | 750.5 | mm of Hg
RoomTemp: | 21.5 | C
#2 | 121.5 | 21 | 1.3402 | 0.045 | 0.04413
| 156.9 | 21 | 1.7446 | 0.065 | 0.063743
| 205.4 | 21 | 2.2986 | 0.12 | 0.11768
| 303.2 | 21 | 3.4157 | 0.28 | 0.27459
| 355.9 | 21 | 4.0177 | 0.4 | 0.39227

## GK 4.162
GK 4.162 Respirable Cyclone
Q-lpm | D50
7 | 4.89
8 | 4.34
8.5 | 4.09
9 | 3.91
9.5 | 3.67
10 | 3.47
Work Area
Enter in RED
Answer in BLUE
Q-lpm | DP
8.5 | 4.0848
GK 4.162
2011-08-23 | Operator: | rg
GK 4.162 with Empty 47 mm Cassette
Qs corrected to: BP=760 mm of Hg, T= 20 C
Room BP: | 758.8 | mm of Hg
RoomTemp: | 21.1 | C
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1A | 141.95 | 20.9 | 1021.5587 | 1.3 | 1.2749
| 182.5 | 20.9 | 1321.3258 | 2 | 1.9613
| 243.7 | 20.9 | 1773.7487 | 3.7 | 3.6285
| 295.2 | 21 | 2155.1967 | 5.4 | 5.2956
| 355.4 | 21 | 2600.3784 | 7.8 | 7.6492
Work Area
Enter in RED
Answer in BLUE
Q-lpm | DP
10 | 4.0212

## SCC 0.732
Q | D50 | SCC 0.732 | rg | 2005-09-10
0.5 | 5.006
1 | 2.239
1.5 | 1.398
2 | 1.001
2.5 | 0.773
3 | 0.625
3.5 | 0.523
4 | 0.448
Work Area
Enter in RED
Answer in Blue
Q-lpm | D50
0.9 | 2.5299
2009-01-22 |  |  |  | Work Area
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Enter in RED
Room BP: | 755.7 | mm of Hg |  | Answer in BLUE
RoomTemp: | 20.2 | C |  | Q-lpm | SP
|  |  |  | 10 | 251.7034
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
Room BP: | 750 | mm of Hg
RoomTemp: | 19.6 | C
#2D | 135.9 | 20.5 | 1.4106 | 4.91 | 4.8151
Room BP: | 755.7 | mm of Hg
RoomTemp: | 20.2 | C
#3D | 241.65 | 20.6 | 0.39026 | 0.48 | 0.47072
| 448.9 | 20.6 | 0.74437 | 1.38 | 1.3533
| 532.7 | 20.6 | 0.88755 | 1.84 | 1.8044

## SCC 1.197
Q | D50 | SCC 1.197 | rg | 2005-09-14
1 | 6.479
2 | 2.897
3 | 1.81
4 | 1.296
5 | 1
6 | 0.809
Work Area
Enter in RED
Answer in Blue
Q-lpm | D50
8 | 0.57948
2009-01-20 |  |  |  | Work Area
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Enter in RED
Room BP: | 750.4 | mm of Hg |  | Answer in BLUE
RoomTemp: | 19.6 | C |  | Q-lpm | SP
|  |  |  | 10 | 48.1883
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
Room BP: | 750 | mm of Hg
RoomTemp: | 19.6 | C
#2D | 135.8 | 20.2 | 1.408 | 0.68 | 0.66685
| 249.5 | 20.2 | 2.6242 | 2.45 | 2.4026
| 347.1 | 20.2 | 3.6682 | 5.1 | 5.0014
| 476 | 20.2 | 5.047 | 10.5 | 10.297

## H-D
Higgins and Dewell Cyclone performance model. |  |  |  |  | 2003-10-17 | rg
Based on information provided by Lee Kenny on 23 Oct 94
Q - lpm | D50 - μm
1 | 10.404
2 | 4.8536
3 | 3.1072
4 | 2.2643
5 | 1.7715
6 | 1.4495
7 | 1.2235
8 | 1.0563
9 | 0.92797
10 | 0.82642
Work Area
Input flow rate in red
Answer returns in Blue
| Q - lpm | D50 - μm
| 2.5 | 3.7972
2008-08-27 | Operator: | rg |  | Work Area
H-D |  |  |  | Enter in RED
Qs corrected to: BP=760 mm of Hg, T= 20 C |  |  |  | Answer in BLUE
Room BP: | 762.1 | mm of Hg |  | Q-lpm | SP
RoomTemp: | 22.5 | C |  | 2.2 | 0.4567
Reading
Critical | Abs. P | Temp. |  | DP | DP
Venturi | Crit. Vent. | Crit. Vent. | Qs | DUT | DUT
Number | mm of Hg | C | Lpm | Cm of H2O | hPa
# 1D | 118.9 | 22.3 | 4.657 | 2.5 | 2.4517
| 175.3 | 22.4 | 6.9405 | 5.6 | 5.4917
| 221.6 | 22.5 | 8.8168 | 10 | 9.8066
Room BP: | 762.1 | mm of Hg
RoomTemp: | 22.5 | C
#2D | 85.9 | 22.2 | 0.86628 | 0.06 | 0.05884
| 146.2 | 22.2 | 1.5054 | 0.245 | 0.24026
| 203.6 | 22.3 | 2.1144 | 0.495 | 0.48543
| 299.85 | 22.5 | 3.137 | 1.11 | 1.0885